# DM SERIES PIV LASER

USER OPERATION MANUAL

Photonics Industries International, Inc.

June 2003

# Laser Operation

## Chiller

A THERMOTEK P310 chiller provides cooling water for the laser system. The chiller MUST be started before the power on the laser power supply is turned ON. Two hoses are used to connect the chiller and the laser head. The output port of the chiller is on the right side after the particle filter, and should be connected to the hose with the blue tape strip (blue/cool). The hose with the red tape strip (red/hot) should be connected to the return port.

Always make sure the two hoses are connected properly before starting up the chiller.

Always make sure the chiller has enough cooling water.

- The chiller water temperature should be set at 21.5°C.

- The chiller water flow rate should be 5.8-6.2 LPM.

- The chiller water resistance should be 0.6-0.7Mohm.

Since the chiller consumes a considerable amount of power (7.5A at 220V), be very careful when arranging the power distribution. Do not put the chiller and the laser power supply on the same power strip, or wall plug unless it has a sufficient amount of capacity.

Please follow the THERMOTEK P310 chiller specifications and manual for detailed installation and operation instructions.

Warning: Use Only De-ionized Water for the Chiller!

## AC Power Considerations

The laser system can be configured to operate at the nominal line voltage between 100 to 120 VAC, or between 200 to 240 VAC. The power inlet is located on the rear panel of the laser driver.

## Ventilation Requirements

The rear panel area of the laser driver needs 2 to 4 inches of clearance for air circulation.

## Rear Panel of the Laser Driver

The rear panel has a multiplex connector for the laser head. It MUST be connected before the laser is turned ON.

Three BNC connectors are on the rear panel.

- External Trigger Input. The signal has to be TTL square waveform. The laser pulse is triggered by the falling edge for the external trigger signal.

- External Gate Input. The TTL signal is used to control the laser pulse generation. When the signal is TTL low, the laser pulse can be generated; otherwise, no laser pulse is generated.

- SYNC Out. The signal is exactly the same as the modulation signal of the Q-Switch. This output can be used as a synchronal signal for user devices.

Three DB9 ports are on the rear panel.

- RS-232 to PC (DB9-Male).

- RS-232 to Chiller (DB9-Male).

- User I/O (DB9-Female) – provides the user input interlock function and laser running signal.

Pin 2 & 3: User input interlock. Short for unlock and disconnect.

Pin 1: 	Laser running indicator. High for running.

Pin 4: 	Laser running indicator GND.

## Power-Up Sequence

Before the driver unit is connected to the AC power source, make sure the cable is properly connected from the laser head and the driver unit.

By connecting the driver unit to an AC power source, part of the system, such as temperature control, starts to work. In case the temperature warning occurs, it is necessary to disconnect the driver unit from the AC power source, and wait at least ten minutes before restarting the system.

With the driver unit connected to the AC power source, set the power switch to “ON” (the light from the switch will be ON) to start the power-up sequence.

During the power-up sequence, all LED lights will light once with a beep, and the welcome screen will display. The laser model and the software version are displayed in the lower two lines of the screen. During this time a self-test is performed to ensure the laser head is in the proper status.  If the driver cannot successfully complete this test, the error message will be displayed.

After this self-test, the driver is configured with the parameters it was in when the power was last shut off, and the Status screen displays.

## Front Panel

Described below are the functions of the keys on the front panel, as shown in Figure 1.

- Power Switch – Switches the AC power to the unit ON or OFF.

- Laser Enable Key-Switch – Switches the laser diode driver to Enable or Disable status.

- Remote LED – Lights to show when the unit is being computer controlled, and to show all the parameters that are not accessible to the keys on the front panel (except the LDD, Shutter and Q-SW).

- Fault LED – Blinks when an error occurs.

- Shutter Key – Turns the shutter ON or OFF. Shutter ON makes the laser available to be output, while Shutter OFF cuts off the laser cavity and stops the laser process.

Warning: Never Turn the Shutter to ON When the Diode Current is Over 10 Amps !

- Q-SW Key – Turns the Q-Switch ON or OFF. Since turning the Q-Switch OFF is dangerous for the laser when the power is high, a Q-SW key locking function is provided to disable the Q-SW key operation. The Q-SW key locking function will be described in instruction for LCD menu operation.

Warning: Never Turn OFF the Q-Switch When the Diode Current is Over 10 Amps !

- LDD Key – Turns the laser diode driver ON or OFF. The laser diode driver can output the current to drive the laser diodes ONLY if both the “Laser Enable” key-switch and the “LDD” keys are turned to “Enable” or “ON”.

Warning: Never Turn the Diode ON When the Diode Current Setting is Over 10 Amps !

- Menu Key – Displays the main Menu screen on LCD.

- Home Key – Displays the Status screen, including the most useful status and parameters on LCD.

- Enter Key (►) – Enters the selected menu, makes the selected parameter to be value changing status, or applies the changed value to the parameter and exits the value changing status. Any parameters in changing status will be blinking.

- Cancel Key (◄) – Exit the value changing status without applying the changed value. The exceptions for changed value applying are the diode current setting and the phase match adjustment. These two parameters are applied immediately after the change has been made.

- Previous (▲) / Next (▼) – Move the cursor to select previous or next changeable parameter. The cursor is the black block, which shows up with the parameter alternately. These two keys are also used to change the value of the selected parameters when it is at value changing status. Key Previous (▲) is used to increase the value, while the Key Next (▼) is used to decrease the value. Single pushing the key will increase or decrease the value for one minimum unit. Holding the key will change the value with exponential speed.

- Knob – Besides the Previous (▲) / Next (▼) keys, the knob can be used to change the parameter value. Clockwise turning increases values, while counter-clockwise turning decreases values.

## LCD Menu Operation

The following figure shows the menu structure of the DM Series PIV laser.

Diode Setting...

Is=XX.XXA Ia=XX.XXA

Diode Temp: XX.X C

Diode Volt: XX.XXV

Imax=38.0A Ts=XX.XC C

Pulse Setting...

Q-SW key: Disable

Freq:Ext Gate:Ext

Pulse Rate:XXXXXXHz

Offtime: XXXXX ns

Twin Pulses...

Twin Pulse Mode: XXX

1st Delay:  XXX us

P.Energy 1st: XX.X%

2nd: XX.X%

Pulse sep:   XXX us

2nd offtime: XXXXXns

Information...

… … …

■ Diode Setting

In this submenu, all setting and measurement readings are about diodes.

“Is” – the diode current setting. It is a changeable parameter. The change to this parameter will be applied immediately as long as the diode is “ON”. The “Enter” key is used to exit the current setting status.

“Ia” – the actual diode current reading. This is a read only parameter.

“Diode Temp” – the actual diode temperature reading. This is a read only parameter.

“Diode Volt” – the actual diode voltage reading. This is a read only parameter.

“Imax” – the diode current upper limit setting. This parameter is changeable only if a certain key combination is pushed.

“Ts” – diode temperature setting. This parameter is changeable only if a certain key combination is pushed.

■ Pulse Setting

In this submenu, all setting and measurement readings are about the Q-Switch. All parameters here are changeable except the “Pulse Rate” in “Ext. Freq” mode.

“Q-SW key” – set the Q-SW key on the front panel to “Enable” or “Disable”. Once “Disable” is set, the Q-SW key and LED will keep the current status and will not be accessible.

“Freq” – the trigger source for Q-Switch laser pulse. The source can be either “Ext” (External) or “Int” (Internal). In “Ext” mode, the laser pulse is triggered by falling edge.

“Gate” – the “Gate” signal source for Q-Switch pulse. The source can be either “Ext” (External) or “Int” (Internal). In “Int” mode, laser pulse is always possible; in “Ext” mode, laser pulse is possible only if the “Gate” input is TTL low.

“Pulse Rate” – the Q-Switch modulation frequency. When the trigger source is “Int”, this parameter is changeable; when the trigger source is “Ext”, this parameter is only displayed for external trigger frequency.

“Offtime” – the offtime of the Q-Switch modulation signal. Offtime defines the duration of the laser process in each pulse period.

Offtime

Trigger

Laser Pulse

■ Twin Pulses

In this submenu, all setting and measurement readings are about diodes. All parameters are changeable.

“Twin Pulse Mode” – When the value for this parameter is OFF, The Q-Switch is working in single pulse mode, i.e. there is only one laser pulse generated for each trigger; when the value is ON, the Q-Switch is working in twin pulse mode, i.e. two laser pulses generated for every other trigger.

“1st Delay” – the delay from the trigger to the first pulse.

“ P.Energy 1st & 2nd” – the pulse energy ratio setting for the two pulses.

“Pulse sep” – the separation time from the first pulse to the second pulse.

“2nd offtime” – the offtime for the second pulse.

■ Information

In this submenu, some information is provided.

## DM Laser

The DM-20 laser is specified to operate up to 22mJ at 1.0KHz. Two typical operating power levels are:

- 11.0mJ at 1.0KHz with the laser diode current approximately 22.0A

- 22.0mJ at 1.0KHz with the laser diode current approximately 30.2A

The recommended repetition rate is 1.0-3.0KHz.

The typical pulse duration at 22.0mJ and 1.0KHz is about 300ns.

Please change the repetition rate and laser diode current under the single-pulse operation. After the current is set, it is free to toggle between twin-pulse and single-pulse operation.